ImmunoBioScience Corp. (IBSC) DATA SHEET

FLUOROSHIELD™ with Para Phenylenediamine (PPD) anti fading Mounting Medium

Catalog number: AR-6514-01 (20 ml) (20ml vial comes with 20mm tip and Grey cap),

AR-6514-02 (100 ml), AR-6514-03 (250 ml), AR-6514-05 (1000 ml)

Description: Fluoroshield with strong antifading agent, 1,4-phenylenediamine (PPD) is an aqueous mounting medium for preserving fluorescence of tissue and cell smears. This unique formula prevents rapid photobleaching of FITC, Texas Red, AMCA, Alexa fluoro 488, Alexa fluoro 594, tetramethyly rhodamine, and Redox.  The fluorescence is retained during prolonged storage at 4oC in the dark. This medium contains phenylenediamine and is not suitable for immunofluorescence of Cy dyes, GFP, mCherry, Phycoerythrin (R-PE), phyocyanin (PC), allophycocyanin (APC) and other fluorochromes proteins.

Refractive Index: 1.366 ± 0.002 (This number applies to this mounting medium in solution. Refractive indexes change when the water solvent evaporates and mounting media dries on slides. We do not have the means to measure the refractive indexes of dry mounting mediums; however we expect the numbers to go higher when dried. The refractive index of water is 1.3330)

May encounter problems with FROZEN BRAIN OR OTHER FROZEN TISSUES WITH LOTS OF FAT

Intended Use: Immunofluorescence, confocal microscopy.

Reagent: Ready to use mounting medium with dark coffee color. This color does not interfere with Immunofluorescence.

Storage: 2-8oC is recommended.  Protect from light, DO NOT FREEZE.

Procedure:

- Bring the vial to room temperature.

- Rinse slide to be mounted with DISTILLED OR DEIONIZED WATER, touch the edges of slide on a paper towel to remove excess water. Place slides on a flat surface away from light.

- Turn the vial upside down and open the dropper to remove any air bubbles.

- Apply 2-3 drops of mounting medium directly on top of the specimen.

- Let stand at room temperature for about 3-5 minutes in the dark.

- Apply cover slip carefully avoiding air bubbles.

- The specimen is ready for visualization under a microscope.

- One can seal the edges of cover slip with nail polish, any organic medium or our Organo mounting medium (cat. # AR-6504). If a coverslip is not sealed air bubbles will appear in few days.

- Method for applying Coverslip: Put 1-2 drops of Mounting medium on the specimen. After 3-4 minutes apply coverslip carefully avoiding air bubbles. Put Kimwipes on the top of coverslip, press gently to remove excess mounting medium. With 200 micropipette add organic mounting medium to seal the edges. Incubate at 37C for one hour in the dark to dry organic mounting medium.

- For storage it is recommended that the slide be stored in the dark at 2-8oC.

Removal of Coverslip: Coverslip can be removed before sealing the edges. Soak slide in warm (37oC) distilled or deionized  water for several minutes. Carefully and slowly move the coverslip. Soak in water for an additional few minutes to remove coverslip. Rinse slide several times with warm water to remove all mounting medium. The slide can be remounted again.

Limitation and warranty: Our warranty is limited to the actual price paid for the product. We are not liable for any property damage, personnel injury, time, effort or economic loss due to our product.

MSDS: This product contains 0.05 % sodium azide as a preservative, appropriate care should be taken in handling. National Institute of Occupational Safety and Health has warning that sodium azide can react with lead, copper, brass or solder in the pluming system and forms hydrazoic acid in acidic condition. Discard with copious amount of water. Avoid skin and eye contact with all laboratory products. Use appropriate lab. gear, lab coat , gloves and  safety glasses. Do not ingest any lab. products. This product is not approved for administration in human or animals.

For research use only; not for use in diagnostic procedures. FOR IN VITRO LABORATORY USE ONLY

“In vitro laboratory products for research”

Fluoroshield is a trademark of ImmunoBioScience corp.

Phone: + 1 425 367 4601; Fax: + 1 425 367 4817; cell (mobile) phone: + 1 425 314 0199

Marketing phone: + 1 650 343 IBSC (4272); e-mail:anitaIBSC@aol.com

Web site: www.immunobioscience.com; E-mail: baderbo@gmail.com

Country of origin: USA